## Cover
|  | Module Name | IriEnvoy
|  | Window | IriEnvoy
|  | Android OS | 11
No. | Version(Check List) | Creater | Content | Tested date | Version(Application) | Notes
1
PROJECT TEST: IriEnvoy

## Summary
Status | Function | Total
Passed | 0 | 0
Failed | 0 | 0
Not Test | 0 | 0

## Test case List
Function Name | Sheet Name | Description | Pre-Condition | Remark
Start Capturing | Start Capturing
Device Settings | Device Settings
Result Image | Device Settings
Capture Settings | Capture Settings
Control Button | Control Button
Interface Testing | Interface Testing

## Start Capturing
Testcase ID | Precondition | Objective | Description | Input Type (Valid/InValid) | Input | Expected result | Actual Result | Status (P/F/NT) | Bug ID | Remark
SC-001 | Hardware set up | Verify that can start capturing in auto capturing mode | 1. Click the “Start” button to activate the camera. Assure the auto capture is checked 2. Position the user’s eye to look directly at the mirror on the camera at a distance of about 20 cm (8 in) in front of the camera. 3. The user should open eyes wide and make sure that the iris is reflected on the mirror attached to the top of the camera. 4. The user must approach the lens slowly to bring the iris into focus. The optimal location is 15.5 cm (6.1 inches) away from the front of the camera mirror. The BLUE LED of the device will BLINK when it has detected its first qualified iris image and started capturing iris images. 5. Once The BLUE LED of the device was BLINK, it is best to remain still or move very slowly toward the camera until The BLUE LED was not BLINK and stop turns on, indicating that it has selected a qualified iris image. | Normal | Distance: 15.5-20 cm Eye position: Centered Lighting: Normal room lighting | Camera activates Blue LED blinks during capture Blue LED stays solid when capture complete System successfully captures qualified iris image
SC-002 |  | Verify that cannot start capturing when out of distance range | 1. Click the “Start” button to activate the camera. 2. Position the user’s eye to look directly at the mirror on the camera at a distance of about 30 cm in front of the camera. 2. Position the user’s eye to look directly at the mirror on the camera at a distance of about 5 cm in front of the camera. | Normal | Distance: 30 cm and 5 cm Eye position: Centered Lighting: Normal room lighting | The BLUE LEB of the device won't blink
SC-003 | Hardware has not been set up | Verify that cannot start when no device selected | 1. Click the “Start” button to activate the camera. | Abnormal |  | Warning user to choose device from the device setting
SC-004 |  | Verify that can start capturing in operator capturing mode | 1. Click the “Start” button to activate the camera. Assure the auto capture is unchecked 2. Position the user’s eye to look directly at the mirror on the camera at a distance of about 20 cm (8 in) in front of the camera. 3. The user should open eyes wide and make sure that the iris is reflected on the mirror attached to the top of the camera. 4. The user must approach the lens slowly to bring the iris into focus. The optimal location is 15.5 cm (6.1 inches) away from the front of the camera mirror. 5. When operator decides good and sharp iris patterns have appeared, presses the “Operate Capture” button. 6. The BLUE LED of the device will BLINK when it has detected its first qualified iris image and started capturing iris images. 5. Once The BLUE LED of the device was BLINK, it is best to remain still or move very slowly toward the camera until The BLUE LED was not BLINK and stop turns on, indicating that it has selected a qualified iris image. | Normal | Distance: 15.5-20 cm Eye position: Centered Lighting: Normal room lighting | Camera activates Blue LED blinks during capture Blue LED stays solid when capture complete System successfully captures qualified iris image
SC-005 |  | Verify that cannot start capturing when at dark | 1. Click the “Start” button to activate the camera. 2. Position the user’s eye to look directly at the mirror on the camera at a distance of about 20 cm (8 in) in front of the camera. 3. The user should open eyes wide and make sure that the iris is reflected on the mirror attached to the top of the camera. 4. The user should go place with less light 5. The user must approach the lens slowly to bring the iris into focus. The optimal location is 15.5 cm (6.1 inches) away from the front of the camera mirror. | Abnormal | Distance: 15.5-20 cm Eye position: Centered Lighting: Dark room lighting | The BLUE LEB of the device won't blink
SC-006 |  | Verify that cannot start capturing when mismatching and blurry image | 1. Click the “Start” button to activate the camera. 2. Position the user’s eye to look directly at the mirror on the camera at a distance of about 20 cm (8 in) in front of the camera. 3. The user should open eyes wide and make sure that the iris is reflected on the mirror attached to the top of the camera. 4. The user approach the lens fastly and repeatly | Abnormal | Distance: 15.5-20 cm Eye position: Centered Lighting: Normal room lighting Speed: fast | The BLUE LEB of the device won't blink
SC-007 | capture mode Timebased (secs) | Verify that can start capturing at capture mode Timebased (secs) | 1. Click the “Start” button to activate the camera. Assure the auto capture is checked 2. Position the user’s eye to look directly at the mirror on the camera at a distance of about 20 cm (8 in) in front of the camera. 3. The user should open eyes wide and make sure that the iris is reflected on the mirror attached to the top of the camera. 4. The user must approach the lens slowly to bring the iris into focus. The optimal location is 15.5 cm (6.1 inches) away from the front of the camera mirror. The BLUE LED of the device will BLINK when it has detected its first qualified iris image and started capturing iris images. 5. Once The BLUE LED of the device was BLINK, keep the camera at the time (seconds) you want to capturing | Normal | Distance: 15.5-20 cm Eye position: Centered Lighting: Normal room lighting Capturing mode: Timebased | The BLUE LEB of the device will blink with selected time
SC-008 | capture mode Framebased (secs) | Verify that can start capturing at capture mode Framebased (secs) | 1. Click the “Start” button to activate the camera. Assure the auto capture is checked 2. Position the user’s eye to look directly at the mirror on the camera at a distance of about 20 cm (8 in) in front of the camera. 3. The user should open eyes wide and make sure that the iris is reflected on the mirror attached to the top of the camera. 4. The user must approach the lens slowly to bring the iris into focus. The optimal location is 15.5 cm (6.1 inches) away from the front of the camera mirror. The BLUE LED of the device will BLINK when it has detected its first qualified iris image and started capturing iris images. 5. Once The BLUE LED of the device was BLINK, keep the camera until it stop BLINKING | Abnormal | Distance: 15.5-20 cm Eye position: Centered Lighting: Normal room lighting Capturing mode: framebased | The device's GREEN LED will flash the set number of times.
SC-009 | Save stream option checked | Verify that can start capturing at save stream mode | 1. Click the “Start” button to activate the camera. Assure the save stream is checked 2. Position the user’s eye to look directly at the mirror on the camera at a distance of about 20 cm (8 in) in front of the camera. 3. The user should open eyes wide and make sure that the iris is reflected on the mirror attached to the top of the camera. 4. The user must approach the lens slowly to bring the iris into focus. The optimal location is 15.5 cm (6.1 inches) away from the front of the camera mirror. The BLUE LED of the device will BLINK when it has detected its first qualified iris image and started capturing iris images. 5. Once The BLUE LED of the device was BLINK, it is best to remain still or move very slowly toward the camera until The BLUE LED was not BLINK and stop turns on, indicating that it has selected a qualified iris image. | normal | Save stream checked | Camera activates Blue LED blinks during capture Blue LED stays solid when capture complete System successfully captures qualified iris image Stream is stored in the same directory with specified stream image format
SC-010 |  | Verify that cannot capturing when Side Gazing & Partial Iris | 1. Click the “Start” button to activate the camera. 2. Change eye direction, make eye partial | Abnormal | Distance: 15.5-20 cm Eye position: Centered Lighting: Normal room lighting Eye direction: gazing, partial | The BLUE LEB of the device won't blink
SC-011 |  | Verify that cannot capturing when wearing reflection glasses | 1. Click the “Start” button to activate the camera. 2. Wear reflection glasses | Abnormal | Distance: 15.5-20 cm Eye position: Centered Lighting: Normal room lighting Wear reflection glasses | The BLUE LEB of the device won't blink
SC-012 |  | Verify that cannot capturing when eye image is occlusion | 1. Click the “Start” button to activate the camera. 2. Make the eye squint | Abnormal | Distance: 15.5-20 cm Eye position: Centered Lighting: Normal room lighting Squint | The BLUE LEB of the device won't blink

## Device Settings
Testcase ID | Precondition | Objective | Description | Input Type (Valid/InValid) | Input | Expected result | Actual Result | Status (P/F/NT) | Bug ID | Remark
DS-001 | Device Setting window is open | Verify USB communication channel selection | Select USB from Communication Channel dropdown | Normal | Select "USB" option | USB should be selected and displayed in dropdown
DS-002 | Device Setting window is open | Verify empty communication channel selection | Try to leave Communication Channel empty | Abbormal | No selection | System should not allow empty selection, USB should remain as default
DS-003 | Device Setting window is open | Verify Scan button functionality | Click Scan button with USB communication channel selected | Normal | Click "Scan" button | System should scan and populate Device List with available devices
DS-004 | Device Setting window is open | Verify device selection from list | Select a device from populated Device List | Normal | Click on a device in the list | Selected device should be highlighted
DS-005 | No devices connected | Verify empty device list handling | Click Scan button with no devices connected | Abnormal | Click "Scan" button | Device List should remain empty
DS-006 | Device is selected from Device List | Verify Device Information display | Check if Device Information field shows correct details | Normal | Select connected device | Device Information field should display relevant device details
DS-007 | Device Setting window is open | Verify Device Information button functionality | Click Device Information button | Normal | Click "Device Information" button | Should display detailed device information in device information field
DS-008 | Device Setting window is open | Verify Browse button functionality | Click Browse button to select firmware file | Normal | Click "Browse" button | File explorer should open to select firmware file
DS-009 | Device connected and firmware file selected | Verify Update Firmware button functionality | Click Update Firmware button with valid firmware file | Normal | Click "Update Firmware" button | Firmware update process should begin with progress indication
DS-010 | Device connected | Verify firmware update with invalid file | Attempt firmware update with invalid file type | Abnormal | Select non-firmware file | System should display error message and prevent update
DS-011 | Device Setting window is open | Verify OK button functionality | Click OK button after device configuration | Normal | Click "OK" button | Window should close and save settings
DS-012 | Device Setting window is open | Verify window close functionality | Click X button in window header | Normal | Click "X" button | Window should close without saving changes
DS-013 | Device Setting window is open | Verify window responsiveness | Resize window using corner drag | Normal | Drag window corner | Window should maintain layout and aspect ratio
DS-014 | Device Setting window is open | Verify field interaction during firmware update | Try to change settings during firmware update | Abnormal | Attempt to change communication channel during update | All controls should be disabled during firmware update
DS-014 | Device Setting window is open | Verify multiple device handling | Connect multiple devices and scan | Normal | Connect multiple devices and click Scan | All connected devices should appear in Device List

## Result Image
Testcase ID | Precondition | Objective | Description | Input Type (Valid/InValid) | Input | Expected result | Actual Result | Status (P/F/NT) | Bug ID | Remark
RI-001 | Capture Demo window is open | Verify Eye Subtype selection | Test Eye Subtype dropdown functionality | Normal | Select "UNDEF EYE" from dropdown | Eye Subtype of result image should be set to "UNDEF EYE"
RI-002 | Capture Demo window is open | Verify Image Header selection | Test different header options (None, 2005, 2011, Iritech) | Normal | Select each radio button option | Selected header type should be applied
RI-003 | Capture Demo window is open | Verify Format selection | Test different format selection: RAW, PNG, JPEG, JPEG2000 | Normal | Select each image format option | Format should be set to image format selected
RI-004 | Capture Demo window is open | Verify Kind selection | Test different image kind selection: K1, K2, K3, K7 | Normal | Select each image kind option | Kind should be set to selected image kind
RI-005 | Capture Demo window is open, image format is not RAW | Verify Compression Criteria selection | Select compression criteria from dropdown: size, quality | Normal | Select each compression criteria option | Compression criteria be set to selected option
RI-006 | Capture Demo window is open | Verify Show Metric Quality checkbox | Toggle Show Metric Quality checkbox | Normal | Click Show Metric Quality checkbox | Metric Quality display should toggle
RI-007 | Capture Demo window is open, image format is RAW | Verify cannot choose Compression Criteria selection | Change image format to RAW | Normal | Image format RAW | Compression criteria is disabled
RI-008 | Capture Demo window is open, compression criteria is size | Verify size compression criteria input | Test compression size with different size values | Normal | Select compression size and enter random valid number value | Compression size should be set to input value
RI-009 | Capture Demo window is open, compression criteria is size | Verify invalud size compression criteria input | Test compression size with invalid size values | Abnormal | Enter negative or non-numeric value,number bigger than 9999999 | System should show notthing in the input
RI-010 | Capture Demo window is open, compression criteria is quality | Verify quality compression criteria input | Test compression quality with different quality values | Normal | Select compression quality and enter random valid number value | Compression quality should be set to input value
RI-011 | Capture Demo window is open, compression criteria is quality | Verify invalid quality compression criteria input | Test compression size with invalid quality values | Abnormal | Enter negative or non-numeric value,number bigger than 100 | System should show notthing in the input

## Capture Settings
Testcase ID | Precondition | Objective | Description | Input Type (Valid/InValid) | Input | Expected result | Actual Result | Status (P/F/NT) | Bug ID | Remark
CS-001 | Capture Demo window is open | Verify Eye Subtype selection | Test Eye Subtype dropdown functionality | Normal | Select "UNDEF EYE" from dropdown | Eye Subtype should be set to "UNDEF EYE"
CS-002 | Capture Demo window is open | Verify Capture Mode settings TimeBased(s) | Test Timebased capture mode with different second values | Normal | Select "Timebased (secs)" and enter "3" | Capture mode should be set to timebased with 3 seconds interval
CS-003 | Capture Demo window is open, capture mode Timebased(s) | Verify invalid Capture Mode time input | Enter invalid time value for Timebased capture | Abnormal | Enter negative or non-numeric value | System should show notthing in the input
CS-004 | Capture Demo window is open | Verify Downscale Factor selection | Test different downscale factor values: 1, 2, 4, 8, 16, 32 | Normal | Select each option from dropdown | Image downscaling should reflect selected factor
CS-005 | Capture Demo window is open | Verify Stream Format selection | Test RAW format selection | Normal | Select "RAW" from dropdown | Stream format should be set to RAW
CS-006 | Capture Demo window is open, stream format is not RAW | Verify Streaming Quality input | Test different streaming quality values | Normal | Enter "100" | Streaming quality should be set to 100
CS-007 | Capture Demo window is open | Verify Frame Rate input | Test frame rate value entry | Normal | Enter "20" | Frame rate should be set to 20 fps
CS-008 | Capture Demo window is open | Verify Auto Capture checkbox | Toggle Auto Capture checkbox | Normal | Click Auto Capture checkbox | Auto Capture should be enabled/disabled
CS-009 | Capture Demo window is open | Verify Save Stream checkbox | Toggle Save Stream checkbox | Normal | Click Save Stream checkbox | Save Stream option should be enabled/disabled
CS-010 | Capture Demo window is open | Verify Capture Mode settings FrameBased | Test FrameBased capture mode with different frame values | Normal | Select "Framebased" and enter "20" | Capture mode should be set to framebased with 20 frames
CS-011 | Capture Demo window is open, capture mode Framebased | Verify invalid Capture Mode frame input | Enter invalid frame value for Framebased capture | Abnormal | Enter negative or non-numeric value | System should show notthing in the input
CS-012 | Capture Demo window is open, stream format is not RAW | Verify invalid Streaming Quality input | Enter invalid Streaming quality value for Streaming quality | Abnormal | Enter negative or non-numeric value,number bigger than 100 | System should show notthing in the input
CS-013 | Capture Demo window is open | Verify invalid frame Rate input | Enter invalid frame rate value f | Abnormal | Enter negative or non-numeric value,number bigger than 30 | System should show notthing in the input

## Control Button
Testcase ID | Precondition | Objective | Description | Input Type (Valid/InValid) | Input | Expected result | Actual Result | Status (P/F/NT) | Bug ID | Remark
CB-001 | Capture Demo window is open | Verify Device Setting button | Test Device Setting button functionality | Normal | Click Device Setting button | Device Setting window should open
CB-002 | Capture Demo window is open | Verify Start button | Test Start button functionality | Normal | Click Start button | Capture process should begin
CB-003 | Capture in progress | Verify Stop button | Test Stop button functionality | Normal | Click Stop button | Capture process should stop
CB-004 | Capture in progress | Verify Force Finish button | Test Force Finish functionality | Normal | Click Force Finish button | Current capture should force finish and find best image
CB-005 | Capture complete | Verify Get button | Test Get button functionality | Normal | Click Get button | Should retrieve captured image with the image settings and show image location
CB-006 | Capture in progress,auto capture uncheck | Verify operate capture button | Test operate capture button functionality | Normal | Click operate capture button | Should begin capturing when user get qualified position
CB-012 | Capture Demo window is open | Verify window close functionality | Click X button in window header | Normal | Click "X" button | Window should close

## Interface Testing
Testcase ID | Precondition | Objective | Description | Input Type (Valid/InValid) | Input | Expected result | Actual Result | Status (P/F/NT) | Bug ID | Remark
G-001 | Device not be injected | Verify that can change eye subtype Chrome, Ms Edge | Eye subtype can be change to UNDEF EYE, LEFT EYE, RIGHT EYE | Normal
G-002 | Capture demo screen | Verify that can change capture mode | Capture mode can be change to Timebased (secs) and Framebased | Normal
G-003 | Capture demo screen,Capture mode Timebased (secs) | Verify that can change seconds | Timebased can be change to a different time value seconds | Normal
G-004 | Capture demo screen, Capture mode Framebased | Verify that can change frames | Framebased can be change to a different time value seconds | Normal
G-005 | Capture demo screen | verify that can change downscale factor | Downscale factor can be can to specific value (1, 2, 4, 8, 16, 32) | Normal
G-006 | Capture demo screen | Verify that can change stream format | Stream format can be change to RAW or JPEG | Normal
G-007 | Capture demo screen, Stream format is JPEG | Verify that can change stream quality | Stream quality is from 1 to 100 | Normal
G-008 | Capture demo screen | Verify that can change frame rate | Frame rate is from 1 to 30 | Normal
G-009 | Capture demo screen, Auto capture option unchecked | Verify that can check auto capture option | Auto capture option can be checked | Normal
G-010 | Capture demo screen, Auto capture option checked | Verify that can uncheck auto capture option | Auto capture option can be unchecked | Normal
G-011 | Capture demo screen, Save stream option unchecked | Verify that can check save stream option | Save stream option can be checked | Normal
G-012 | Capture demo screen, Save stream option checked | Verify that can uncheck save stream option | Save stream option can be unchecked | Normal
G-013 | Capture demo screen | Verify that can change image header | Image header can be none, 2005, 2011, Iritech i image header is disabled | Normal
G-014 | Capture demo screen | Veirfy that can change change result image format | The result image format can be RAW, JPEG, PNG, JPEG2000 | Normal
G-015 | Capture demo screen | Verify that can change kind of result image | The result image kind can be K1, K2, K3, K7 | Normal
G-016 | Capture demo screen, Result image format is PNG, JPEG, JPEG2000 | Verify that can change compression criteria of result image | The compression criteria can be size or quality | Normal
G-017 | Capture demo screen, Compression criteria is Size | Verìy that can change the compression size | The compression size is from 0 to 9999999 | Normal
G-018 | Capture demo screen, Compression criteria is Quality | Verify that can change the compression quality | The compression size is from 0 to 9999999 | Normal
G-019 | Capture demo screen, Show metric quality option unchecked | Verify that can check show metric quality option | Show metric quality can be checked | Normal
G-020 | Capture demo screen, Show metric quality option checked | Verify that can uncheck show metric quality option | Show metric quality can be unchecked | Normal
G-021 | Capture demo screen | Verify that can click device setting button | Click the device setting button to show device setting screen | Normal
G-022 | Device setting sceen | Verify that can change the communication channel | change the communication channel | Normal
G-023 | Device setting sceen | Verify that can click scan device button | Click scan | Normal
G-024 | Device setting sceen | Verify that can change device list | Change device option | Normal
G-025 | Device setting sceen | Verify that can click browse button | Cick brown | Normal
G-026 | Device setting sceen | Verify that can click device information button | Click device infomation | Normal
G-027 | Device setting sceen | Verify that can click update firmware button | Click update firmware | Normal
G-028 | Device setting sceen | Verify that can click ok button | Click ok | Normal
G-029 | Device setting sceen | Verify that can click get button | Click get | Normal